A. General

1. This is not a complete uncertainty calculation.

Only those contributions to the uncertainty of the efficiency are estimated,

which are coming from of the different nuclear and atomic data.

2. All contributions are calculated with the entry parameters which are given in the head of this protocol:

kB value

the number of Monte Carlo events

the reference value of fm.

3. Other sources of uncertainties, which are more connected with the measurement, are

| T1/2 | The date of the measurement determines the uncertainty. Calculate this contribution with the evaluation of measurements. |
| --- | --- |
| sample stability | Calculate the contribution of the sample preparation with the evaluation of measurements. |
| impurities | The influence of impurities is to calculate individually. |
| tracer activity | Calculate the contribution of the uncertainty of the activity of the tracer standard with the evaluation of measurements |
| tracer calibration | Calculate the contribution of the goodness of the tracer calibration with the evaluation of measurements |
| KB value | An influence of the kB value is given with an efficiency calculation with more than one kB value. The kB value depends on the composition of the samples and is not well known. |

4. The calculation of the uncertainty due to nuclide and atomic data is done in the following way:

# Option Actual Uncertainty

If  uncertainties are known then the data and their uncertainties are read in from the file CNUnc.nuc.

This file can be supplemented with data sets of new nuclides.

# Option Uncertainty Preview

If the uncertainties are unknown then typical uncertainties are read in from the file CNGenUncert.in.

The input data are changed with typical uncertainties in a realistic way: All uncertainties are added or

subtracted to the input data in such a direction, that each individual change results a lower efficiency.

Then the efficiency of the part of the decay is calculated with changed data and with the original data. The difference between both is set as the standard deviation of the decay part.

The influence of all uncertainty compartments is followed to estimate the nuclide uncertainty.

Details: see protocol.