user  manual

for the TDCRB-1 computing program

(version 1.1)

R. Broda1*,		rbroda@cyf.gov.pl

P. Cassette2,		Philippe.Cassette@cea.fr

K. Małetka1,		e08km@cx1.cyf.gov.pl

K. Pochwalski3	tofer@it.com.pl

1 : RC POLATOM, 05-400 Otwock-Swierk, Poland

2 : CEA/DAMRI/LPRI, BP 52, 91193 Gif sur Yvette Cedex, France

3 : retired from RC POLATOM

* Who is responsible for programming

Swierk, Saclay, 1999

__________________   *  __________________

The User Manual is related to the paper:

R. Broda, P. Cassette, K. Małetka, K. Pochwalski, "A simple computing program for application of the TDCR method to standardization of pure-beta emitters", submitted to the ICRM'99, Prague, Czech Republic, 1999.

Contents

1.	Definitions and notions ...........................................................2

2.	Input data .................................................................................3

3.	Program operation ...................................................................6

4.	The program structure .............................................................9

5.	Output data ............................................................................12

The TDCRB-1 program in FORTRAN language have been developed to calculate the detection efficiency and hence the source activity of a triple photomultiplier LS counter for pure beta-emitting radionuclides. The program is based on the TDCR method data processing.

1.	Definitions of notions

Notions and quantities used in the program are defined as follows:

| Notions | Definitions |
| --- | --- |
| Source Set of sources Source mass Source activity Solution activity Counting time Count Repetition Counting channel Counting rate Figure of merit Counting efficiency Counting point Experimental TDCR parameter Theoretical TDCR parameter Set of counting points Counting point activity Measurement Results Information concerning the measurement Input data set | a vial filled with a cocktail of scintillation solution with the radioactive solution to be measured all sources prepared from the radioactive solution to be measured amount of the radioactive solution in the scintillation cocktail of a source [mg] radioactive concentration of the radioactive solution contained in the source [kBq/g] mean radioactive concentration of the solution contained in the set of sources measured [kBq/g] a basic length of time set to count pulses [s] a number of pulses counted in the set counting time (refers to: a radioactive source or a background source, and also to a source with background subtracted) a set number of successively taken counts (in the same counting time) one of AB, BC, CA, T or D coincidence channels, in which counts are taken a number of pulses counted per unit time in a given counting channel [cps] a parameter describing actual quality of a scintillation detector, defined as a number of photoelectrons collected at first dynodes of all PM-tubes, per incident electron energy deposited in the scintillator [photoelectrons/keV] characterized by a figure of merit, theoretically calculated counting probability of scintillation pulses, integrated over the whole energy spectrum a set of 5 means of successively repeated counting rates in all 5 counting channels at a single given figure of merit (interrelated to the experimental TDCR parameter) [cps] ratio of counting points in T and D channel ratio of counting efficiencies in T and D channel of the corresponding counting point all counting points successively taken at a set number of different TDCR parameter (different figure of merit) values source activity corresponding to a given counting point [kBq/g] a final set of measurement results data general text description of the measurement and conditions applied a set of input numerical data of counts and parameters to be processed to obtain measurement results |

2.	Input data

An example of a measurement results input data file NI6301.TDK, attached to the manual, is given. The file contains counting results of a set of three 63Ni sources in the Ultima Gold scintillator, in 7 counting points at different efficiency, which have been obtained by changing the PM-tubes focalisation, and measurement in each counting point was repeated 3 times, 60 s each). The background data file BKGR3.TDK, attached to the manual, is of the same format.

Underlined variables serve only as an information concerning measurements and after reading are not processed. The integer (I), real (R) and double precision (DP) variables are read by using a free format. Dimension of the character (CH) variables are given in a table.

Names of radionuclide or scintillator in an input data file DONNE.TAB, attached to the manual, can be changed by a user, but numbers JB or JS should be actualized in the same time.

| Format of the input data set file | Name of variables used | Type |
| --- | --- | --- |
| Ni6301.tdk Ni-63 23/ 2/1996 Ultima Gold (EUROMET) defocalisation 3 7 46.055 47.082 52.921 800 775 750 725 700 675 650 60 3 81682.0 82332.0 80267.0 76317.0 91647.0 81336.0 81970.0 79730.0 75857.0 91322.0 81295.0 81956.0 79936.0 76022.0 91143.0 80910.0 81738.0 79462.0 75533.0 91044.0 81604.0 82315.0 80100.0 76167.0 91685.0 81643.0 82116.0 79972.0 76098.0 91535.0 ... ... ... ... ... 71066.0 74408.0 70503.0 65422.0 85133.0 69998.0 73755.0 69619.0 64488.0 84396.0 69478.0 73485.0 69138.0 63995.0 84111.0 60 3 83612.0 84192.0 81782.0 77934.0 93718.0 83281.0 83971.0 81366.0 77617.0 93384.0 82590.0 83511.0 80860.0 77154.0 92653.0 82596.0 83163.0 80641.0 76895.0 92610.0 82328.0 83079.0 80367.0 76667.0 92440.0 82990.0 83855.0 81069.0 77308.0 93298.0 82129.0 83266.0 80368.0 76542.0 92679.0 82317.0 83598.0 80620.0 76796.0 92943.0 ... ... ... ... ... | NOMFIC NUCL DATE INFS NS NP MASS(i) for i=1, ... , NS UFOC(j) for j=1, ... , NP T(1) NR(1) counting rate of coincidence channels AB, BC, CA, T, D as X(1,j), ... , X(5,j) for j=1, ... , NR (NP counting points of the 1st source) T(2) NR(2) X(1,j), ... , X(5,j) for j=1, ... , NR (NP counting points of the 2nd source) | CH*14 CH*9 CH*12 CH*80 I I R R R I DP R I DP |
| Meaning of the variables used: NOMFIC - name of the input data set file NUCL - name of the radionuclide DATE - date of taking measurement INFS - information concerning the measurement NS - number of sources NP - number of counting points MASS - table of source masses [mg] | UFOC - table of voltage values applied to the focusing electrode T - table of counting times applied to different sources NR - table of repetitions applied to different sources X - table of all repetitions of all counting points of different sources |  |

| Format of the DONNE.TAB file | Name of variables used | Type |
| --- | --- | --- |
| 8 NI-63 63. 28. -1 65.870 H-3 3. 1. -1 18.620 TL-204 204. 81. -2 763.400 C-14 14. 6. -1 156.478 SR-90 90. 38. -2 546.000 Y-90 90. 39. -2 2284.000 W-188 188. 74. -1 349.000 RE-188 188. 75. -1 2118.000 3 Toluene + (common carrier solution); 8.3ml 1.15 2. .867 .54271 .06103 Toluene + (Ni+alk.+TOPO; evaporated); 10.0ml 1.27 1.97 .867 .54271 .06103 Ultima-Gold /according to PTB/; 10.6 ml 1.27 2.08 .985 .546627 .065115 | JB NUCL NM NA IFORBN EMAX 2nd radionuclide 3rd radionuclide 4th radionuclide 5th radionuclide 6th radionuclide 7th radionuclide 8th radionuclide JS DAN R H RO ZA EJ 2nd scintillator 3rd scintillator | I CH*9 R R I R I CH*70 R R R R R |
| Meaning of variables used: JB - number of different radionuclides in the file NUCL - name of a radionuclide NM - the radionuclide mass number NA - the radionuclide atomic number IFORBN- forbiddenes of the beta transition (an absolute value of the nuclear spin change; sign corresponding to the character of the beta decay - beta minus or plus) EMAX - maximal energy of the beta spectrum [keV] | JS - number of different scintillators in the file DAN - name of a scintillator R - radius of the scintillator cylindrical column [cm] H - height of the scintillator cylindrical column [cm] RO - density of the scintillator [g/cm3] ZA - ratio of the atomic to the mass number of the scintillator EJ - mean exciting energy of the scintillator molecules [keV] |  |

3.	Program operation

The TDCRB-1 program communicates with the user in an interactive way. Several successive questions are displayed and optional alphanumerical answers are proposed in brackets to select. The program can be used in two counting modes, depending on the variable IEFF value selected. Examples of results of calculations displayed are given in Chapter 5.

Selection of the way of presentation of results:

==>	Print the results of calculation 	(1)

or save them in a file CALC.RES ?	(2)

1 - Calculation results are directly printed (this option is useful when continuous printer paper isused);  2 - (recommended) results are saved in a new version of CALC.RES file.

Selection of the calculation mode (IEFF value):

==>	Calculation of the solution radioactive concentration  (1)

or the theoretical counting efficiency only ?          (2)

1 - calculation of a solution activity, based on the input data set file (the IEFF variable takes the value of 1);

2 - calculation of a theoretical counting efficiency for a given set of thirty TDCR parameter values (the IEFF variable takes the value of 2).

3.1.	IEFF=1 option

This part of the program appears only, if the option IEFF=1 is selected.

==>	Do you want to read the file with background counts ?  (Y/N)

The calculation is possible with (“Y”), or without taken into account the background data file (“N”).

==>	Name of the input data set file with a background counts:

The full name of the background data file saved on the computer disc should be input.

***	SUCH A FILE DO NOT EXIST !!!

==>	Do you want to STOP the program ? (Y/N)

This warning appears, if the file name has not been correctly input.

==>	Name of the input data set file with a set of sources counts:

The full name of the measurement results input data file saved on the computer disc should beinput.

***  SUCH A FILE DO NOT EXIST !!!

==>	Do you want to STOP the program ? (Y/N)

This warning appears if the file name has not been correctly input.

3.2.	Both IEFF=1 and IEFF=2 options

This part of the program appears independently of the IEFF variable value.

==>	Select the radionuclide for calculation (by “Y”)

Radionuclide symbols list from the DONNE.TAB file appears on the monitor one by one. Pressing any key except of “Y” displays the next symbol. Pressing “Y” means selection of a given radionuclide After checking the list of radionuclides without pressing “Y” the warning is displayed:

==>	Do you want to STOP the program ? (Y/N)

When any key except “Y” is pressed, the list of radionuclide symbols is displayed from the beginning.

==>	Select the scintillator (by “Y”):

Scintillators names list from the DONNE.TAB file appears on the monitor one by one. Pressing any key except of “Y” displays next scintillator name. Pressing “Y” means selection of the scintillator used in the measured source. After checking the list of radionuclides without pressing “Y” the warning is displayed:

==>	Do you want to STOP the program ? (Y/N)

When any key except of “Y” is pressed, the list of the scintillators names is displayed from the beginning.

Following values of the kB parameter are chosen for interpolation:

.0060    .0090    .0120    .0150    .0180

==>	Do you accept above values ? (Y/N)

By pressing any key except “N”, you accept an interpolation with a given set of the kB parameter values. By pressing “N”, you can introduce another set of the kB parameter values. Then you must input five kB parameter values following the instruction, e.g.:

==>	A new value of kB(1) = .0015

==>	A new value of kB(2) = .003

==>	A new value of kB(3) = .005

==>	A new value of kB(4) = .007

==>	A new value of kB(5) = .009

and the question of accepting a set of the kB parameter values appears again.

3.3.	IEFF=1 option

This part of program appears only, if the option IEFF=1 was selected.

==>	Select the printing option:

- tables and diagrams for all kB values			(1)

- diagrams but not tables for all kB values		(2)

- table and diagram only for interpolated kB value	(3)

1 - Full calculation results are printed.

2 - Diagrams of a set of counting points for each source of a set, for each kB parameter

value and the final table of results of calculation of a set of sources are printed.

3 - Diagrams of a set of counting points for each source of a set, for interpolated kB

parameter value only and the final table of results of calculation of a set of sources are printed.

==>	Calculation of the theoretical counting efficiency table

for the mean value of the interpolated kB parameter ? (Y/N)

Pressing “N” ends the program.

When any key except “N” is pressed,

ETA0 - figure of merit, PHEL - expected number of photoelectrons in total, effC - counting efficiency of double coincidence, effT - counting efficiency of triple coincidence, effD - counting efficiency of sum of double coincidence, for 30 different values of the TDCR ratio in an interval [0.02, 0.99] are calculated.

The mean value of the interpolated kB parameter is taken into calculation in each case.

==>	Calculation for an other TDCR value ? (Y/N)

Pressing “N” ends the program. When any key except “N” is pressed,  the following instruction is dispalyed:

==>	Give a new value of TDCR =

When the TDCR parameter value is input, calculation of ETA0, PHEL, effC, effT and effD values for a given TDCR parameter value and for the same as before kB parameter valueis performed.

3.4.	IEFF=2 option

This part of the program appears only, if the option IEFF=2 was selected.

==>	Calculation for another TDCR value ? (Y/N)

Pressing “N” ends the program. When any key except of “N” is pressed, the following instruction is displayed:

==>	Give a new value of TDCR =

When the TDCR parameter value is input, calculation of ETA0, PHEL, effC, effT and effD values for a given TDCR parameter value and for the set of the kB parameter values is preformed.

4.	The program structure

The main TDCRB-1 program and a set of subroutines have been compiled independently. The program may be used in two operating modes, depending on the IEFF parameter value selected (see Chapter 3 and next two diagrams). The program TDCRB-1 accepts an input data set file with up to: 30 counting points, 15 repetitions and 12 sources. The radionuclide and the scintillator data are selected from the DONNE.TAB file. Calculation results are printed or saved in the CALC.RES file on a computer disc.

Subroutines:

DATAEXP	- For selecting the input data set file with a set of sources, as well as the

input data set file with a background counts. Means of counts are taken, counting points are calculated, background is subtracted and uncertainties calculated.

PARAM	- For selecting and printing the radionuclide and the scintillator data,

registered in the DONNE.TAB file. The data are used to calculate the electron and scintillation beta spectra.

SPECTB	- For calculation beta-spectrum of a given radionuclide. The spectrum

is divided into a selected number of intervals.

FJON		- For calculating scintillation energies, using the kB parameter and

selected scintillator of a given dimensions, in all intervals of the beta spectrum.

FIT			- For calculating counting efficiencies at all counting channels of a

counting point, by fitting the calculated theoretical TDCR parameter (Kth) value to the experimental one (Kexp).

SOURCE	- For calculating activity of the source as a mean of all counting point

activities. For this purpose a slope of a regression line of the set of counting points is calculated.

DESSIN		- For printing tables and diagrams of a set of counting points as a

function of the TDCR parameter.

INTRP		- For interpolation of function values between given points by using the

Lagrange polynomial of N-1 degree.

SETSOUR	- For calculating the mean activity of the solution.

RESULT	- For printing results of calculation.

The first mode (IEFF=1) of the TDCRB-1 program operation

(calculation of a solution radioactive concentration):

RESULT

SETSOUR

STOP (1)

solution results

NO

YES

NO

diagram of

counting points

DESSIN

YES

INTRP

kB

interpolation

done ?

last

kB value

?

SOURCE

YES

NO

last

counting point

?

counting point

activity

FIT

next

counting point

FJON

next kB

value

next source

SPECTB

PARAM

experimental

parameter TDCR

DATAEXP

START (1)

parameters

used

radionuclide and

scintillator data

file DONNE

counting data

file

YES

last

source

?

NO

solution

activity

best

kB value

source

activity

counting

efficiency

scintillation

spectrum

beta

spectrum

The second mode (IEFF=2) of the TDCRB-1 program operation

(calculation of a theoretical efficiency for a given set of TDCR parameter values)

new TDCR value

NO

last

kB value

?

YES

NO

STOP (2)

YES

other

TDCR value

?

last

TDCR value

?

counting

efficiency

next TDCR

value

FIT

NO

YES

START (2)

set of TDCR

parameters

PARAM

parameters

used

radionuclide and

scintillator data

file DONNE

SPECTB

next kB value

FJON

counting

efficiency

scintillation

spectrum

beta

spectrum

5.  Output data

5.1.	IEFF=1 option

An example of a part of results printout (the first mode of the program operation; IEFF=1) for the set of 63Ni sources measurement input file Ni6301.TDK and the background data file BKGR3.TDK

The table of single source calculation results, obtained with an interpolated kB parameter value and a diagram of a set of counting points is given:

a)

A      - counting point activity [kBq/g]

SIGMA  - counting point activity uncertainty (standard deviation)

ETA0   - figure of merit [photoelectrons/keV]

PHEL   - number of photoelectrons in the detector

DIFF   - test of mutual relation of the counting channels

TDCR   - experimental TDCR parameter

AB, BC, CA, T, D - mean counting rate in a given counting channel

per source mass [cps/mg]

effC, effT, effD - counting efficiencies of counting channels

(assumption: efficiencies effAB=effBC=effCA =effC)

SOURCE No. 1    MASS=  46.0550 mg   COUNTING TIME=  60. s   REPETITIONS= 3

*  kB=  .00788 cm/MeV

A   SIGMA   ETA0        T   effT        D   effD   TDCR

1    39.675    .064   .920    27.42  .6910    32.90  .8294  .8332

2    39.778    .085   .909    27.38  .6883    32.93  .8278  .8315

3    39.689    .154   .878    27.01  .6805    32.67  .8232  .8266

4    39.669    .085   .833    26.51  .6684    32.37  .8161  .8190

5    39.714    .073   .767    25.77  .6488    31.95  .8045  .8065

6    39.628    .155   .693    24.69  .6230    31.27  .7891  .7896

7    39.746    .144   .604    23.30  .5863    30.47  .7666  .7649

PHEL    DIFF       AB       BC       CA   effC

1  12.752  .00000    29.34    29.58    28.82  .7371

2  12.599  .00000    29.33    29.58    28.78  .7348

3  12.168  .00000    28.97    29.30    28.42  .7281

4  11.544  .00000    28.46    29.00    27.94  .7176

5  10.640  .00000    27.71    28.50    27.27  .7007

6   9.604  .00000    26.63    27.71    26.31  .6784

7   8.374  .00000    25.30    26.63    25.14  .6464

A [kBq/g]                            ACT =   39.700 +-  .019 kBq/g

39.87-!

!!                 !

!!               ! !

!!               ! !

!!          !    ! !

39.76-!      !   !    ! 2

!7      !   !  ! ! !

!!      !   !  ! ! !

!!      !   5  ! ! !                                          !

!!      !   !  ! 3 !                                          *

39.66-!      !   !  4 ! 1                                          !

!!      !   !  ! ! !

!!      6      ! ! !

!!      !      ! ! !

!       !      ! !

39.56-       !        !

!       !        !

!       !        !

!       !

!       !

39.45-

!--------------!--------------!--------------!--------------!

.765           .824           .882           .941          1.000 TDCR

The table of a single source calculation results, obtained with different kB parameter values is given below. The last kB value is found by interpolation method, whena zero slope value of a regression line of a set of counting points was taken as a condition.

b)

* FITTING OF A THEORETICAL COUNTING EFFICIENCY TO THE COUNTING POINTS

* BY CHANGING kB PARAMETER VALUE

* Source no. 1

* kB=  .00600 cm/MeV    ACT=   39.44217 +-  .02165 kBq/g    Slope=   2.79 %

* kB=  .00900 cm/MeV    ACT=   39.83662 +-  .01987 kBq/g    Slope=  -1.47 %

* kB=  .01200 cm/MeV    ACT=   40.15495 +-  .02722 kBq/g    Slope=  -4.88 %

* kB=  .01500 cm/MeV    ACT=   40.42072 +-  .03696 kBq/g   Slope=  -7.70 %

* kB=  .01800 cm/MeV    ACT=   40.64785 +-  .04672 kBq/g    Slope=-10.10 %

* kB=  .00788 cm/MeV    ACT=   39.70016 +-  .01906 kBq/g    Slope=     .00 %

The table of final results of a set of sources, an arithmetic mean of radioactive concentration of a solution with their uncertainties, as well as conclusions drawn from the results:

c)

Source   Mass       Activity             kB    Slope

no.        [mg]       [kBq/g]              [cm/MeV]

1      46.06     39.7002 +-  .0191   (+-  .05 %).00788       . .00 %

2      47.08     39.2880 +-  .0332   (+-  .08 %)      .00506      -.01 %

3      52.92     39.1479 +-  .0269   (+-  .07 %)      .00538        .01 %

MEAN              39.3787 +-  .1658   (+-  .42 %)      .00610

==========================================================================

MEAN SOLUTION ACTIVITY =    39.379 +-  .166 kBq/g   ( +-  .42 % )

==========================================================================

CONCLUSIONS  which may be drawn from the results :

* Difference between a source activity No. 3 and No. 1 is   1.39 %

* Difference between kB values for a source No. 2 and No. 1 is  35.85 %

The differences between the highest and the lowest source activity in a set, as well as the differences between the highest and the lowest kB parameter value are calculated and presented in the CONCLUSIONS.

An example below shows the conclusions drawn by the TDCRB-1 program from the results of a set of nine sources with intentionally incorrect input parameters taken to the calculation:

d)

Source   Mass       Activity                              kB            Slope

no.    [mg]       [kBq/g]                          [cm/MeV]

1      16.35     38.6509 +-  .0166   (+-  .04 %)        .03816          2.56 %

2      16.35     39.7400 +-  .0176   (+-  .04 %)        .07764        -2.35 %

3      16.35     39.0538 +-  .0145   (+-  .04 %)        .05013         2.40 %

4      16.35     33.9669 +-  .0952   (+-  .28 %)        .00000       23.14 %

5      16.35     42.5048 +-  .1254   (+-  .30 %)     1.13933    -17.41 %

6      16.35     33.0033 +-  .1350   (+-  .41 %)       .00000       31.61 %

7      16.35     33.1836 +-  .1037   (+-  .31 %)       .00000       22.89 %

8      16.35     39.0425 +-  .0214   (+-  .05 %)       .15816       -1.34 %

9      16.35     44.0580 +-  .1393   (+-  .32 %)     2.09400    -18.15 %

MEAN              38.1338 +- 1.3276   (+- 3.48 %)        .39527

==========================================================================

MEAN SOLUTION ACTIVITY =    38.134 +- 1.328 kBq/g   ( +- 3.48 % )

==========================================================================

CONCLUSIONS  which may be drawn from the results :

* Difference between a source activity No. 6 and No. 9 is  25.09 %

* SPREAD OF THE RESULTS SEEMS TO BE TOO HIGH

* Difference between kB values for a source No. 4 and No. 9 is 100.00 %

* THE DIFFERENCE SEEMS TO BE TOO HIGH

* kB VALUE SEEMS TO BE TOO LOW

* for the source No. 4,  6,  7,

* kB VALUE SEEMS TO BE TOO HIGH

* for the source No. 1,  2,  3,  5,  8,  9,

* A REGRESSION LINE SEEMS TO BE TOO MUCH INCLINED

* for the source No. 1,  2,  3,  4,  5,  6,  7,  9,

Warnings visible in CONCLUSIONS (introduced into the RESULT subroutine) appear if :

- the difference between the highest and the lowest source activity in a set of sources is higher than 3%

- the differences between the highest and the lowest extrapolated kB parameter value in a set of sources is higher than 40%

- the extrapolated kB parameter value is lower than 0.002 cm/MeV or higher than 0.020 cm/MeV

- the slope value of the regression line of a set of counting points for any source is higher than ± 2%

5.2.	IEFF=2 option

An example of printout results for 63Ni (the second mode of the program operation; IEFF=2).

The theoretical counting efficiencies of counting channels, calculated for the first value of ionization quenching kB parameter from a set and different TDCR ratios:

e)

kB     - ionization quenching kB parameter [cm/MeV]

TDCR   - given (experimental) TDCR parameter

ETA0   - figure of merit [photoelectrons/keV]

PHEL   - number of photoelectrons in total in the detector

effC, effT, effD - counting efficiencies of counting channels

THEORETICAL COUNTING EFFICIENCY FOR A GIVEN kB AND TDCR PARAMETER VALUES

kB     TDCR     ETA0     PHEL     effC     effT     effD

.0060    .0200     .006      .09    .0013    .0001    .0037

.0060    .0500     .015      .22    .0075    .0010    .0204

.0060    .0900     .027      .39    .0220    .0050    .0560

.0060    .1200     .037      .53    .0367    .0106    .0887

.0060    .1500     .046      .67    .0538    .0186    .1243

.0060    .1900     .060      .86    .0799    .0330    .1738

.0060    .2200     .070     1.01    .1015    .0465    .2114

.0060    .2500     .081     1.17    .1244    .0622    .2487

.0060    .2900     .097     1.40    .1568    .0863    .2977

.0060    .3200     .109     1.58    .1823    .1067    .3334

.0060    .3500     .122     1.77    .2086    .1289    .3682

.0060    .3900     .141     2.04    .2451    .1611    .4131

.0060    .4200     .157     2.27    .2733    .1872    .4456

.0060    .4500     .174     2.51    .3023    .2148    .4773

.0060    .4900     .199     2.87    .3419    .2539    .5181

.0060    .5200     .220     3.18    .3725    .2848    .5478

.0060    .5500     .244     3.52    .4038    .3172    .5768

The theoretical counting efficiencies of counting channels calculated for selected TDCR ratio and for set of the kB parameter value.

f)

THEORETICAL COUNTING EFFICIENCY FOR A GIVEN kB AND TDCR PARAMETER VALUES

kB     TDCR     ETA0     PHEL     effC     effT     effD

.0060    .5550     .248     3.58    .4090    .3228    .5816

.0090    .5550     .260     3.53    .4016    .3169    .5710

.0120    .5550     .273     3.49    .3955    .3121    .5623

.0150    .5550     .284     3.45    .3904    .3081    .5550

.0180    .5550     .296     3.42    .3860    .3046    .5488